PRODUCT RECALL
IDS-24-5142
BD BBL Sensi Disc Ampicillin

2 μg (AM-2)

07 November 2024,

Dear BD Distributor,

Type of Field Action: Recall

Affected Product

| Product Name (Brand Name as per labelling) | Catalog Number | UDI-DI | Lot No. | Expiration Date |
| --- | --- | --- | --- | --- |
| BD BBL Sensi Disc Ampicillin 2 μg (AM-2) | 231263 | (01)30382902312636 | 2339360 | 2024/12/31 |
|  |  |  | 3010977 | 2025/01/31 |
|  |  |  | 3058508 | 2025/03/31 |
|  |  |  | 3184064 | 2025/07/02 |
|  |  |  | 3234190 | 2025/08/21 |

Description of the Problem:

BD identified through potency test, as part of special stability test requests to monitor Ampicillin AM-2, unsatisfactory results of 65% at 18 months. The potency acceptance criteria is 80-140% and product shelf life is 24 months. A retest with control batch and extended sampling confirmed the low potency results.

Clinical Risk Statement:

A decrease in the potency of ampicillin has been observed in the affected products. This could cause reduced antibiotic diffusion through the disc and thus produce a decreased zone of inhibition. If an organism is truly susceptible to ampicillin, growth may not be inhibited appropriately and interpreted erroneously as resistant. A falsely resistant result for ampicillin susceptibility would unnecessarily exclude ampicillin as a treatment option. In order to mitigate this risk, customers should discard any affected product.

Complaint & Adverse Event Statement:

To date, there have been no complaints and no adverse events worldwide related to this issue.

Actions for Clinical Users:

- Immediately inspect your inventory for the specific catalog and lot number listed in the affected product table above. Destroy affected product subject to the recall following your institution’s process of destruction.

- Ensure the contents of this notification are read and understood.

- There are no additional recommendations for repeat testing or review of prior results.

Action Taken by BD:

BD is investigating the root cause and will implement appropriate corrective actions to prevent recurrence of this issue.

Please Take the Following Actions:

- Immediately inspect your inventory for the specific catalog and lot numbers listed above. Quarantine and destroy affected product subject to the recall following your institution’s process of destruction.

- Disseminate this notice with any users of the product within your facilities or with any interfacility users where product was transferred, to ensure they are also aware of this recall notice.

- Complete and return the attached Distributor Response Form even if you no longer have any inventory remaining in your facility so that BD may acknowledge your receipt of this notification.

- Disseminate this recall notice to all the impacted customers under your distribution. Coordinate customers product return and destruction following your institution’s process of destruction.

- Return the signed and completed Distributor Response Form with Distributor Overview, as well as the signed Customer Response Form from all the impacted customers to the BD contact noted on the form.

- Report any adverse health consequences experienced with the use of these lots to BD.

BD is committed to advancing the world of health. Our primary objectives are patient and user safety and providing you with quality products. We apologize for any inconvenience this issue may have caused you and thank you in advance for helping us to resolve this matter as quickly and effectively as possible.

Yours Sincerely,

_______________________________ Karena Han Quality Manager, Southeast Asia

DISTRIBUTOR RESPONSE FORM
IDS-24-5142

BD BBL Sensi Disc Ampicillin

2 μg (AM-2)

Please fill in the information below so that we may acknowledge your receipt of this notification. 
Complete and return the completed form to SEA_Quality SEA_Quality@bd.com / local BD representative by 19 November 2024.

Please tick as appropriate.

|  | I have read and understood the attached notice taken appropriate actions |
| --- | --- |
|  | We do not have affected product(s) in our inventory. |
|  | We have identified the affected product(s) in our inventory. Affected products have been quarantined until disposal. Upon disposal, we will provide a copy of the Certificate of Destruction/proof of destruction to BD. The expected date of destruction is: ____________________________. |
|  | We have identified all customers that purchased the affected catalog numbers and will notify the affected customers of this notice. The overview of the distribution to the customers are as attached in the Distribution Overview. |

| Product Description | Catalog No. | Lot No. | Quantity Received (EA) | Quantity Sold (EA) | Remaining Quantity (EA) in inventory to be *destroyed |
| --- | --- | --- | --- | --- | --- |

*	Please provide a copy of the Destruction Certificate/ Proof of Destruction for the disposal of all affected units.

Completed by:

| Name: |  |
| --- | --- |
| Signature: |  |
| Date: |  |
| Facility / Address / Telephone Number: |  |