PRODUCT RECALL 
IDS-25-5245
BD BBL™ TB Stain Kit K

14 February 2025

Dear Customer,

CC: Chairman Medical Board and relevant head of department

Type of Field Action:

Removal (Recall)

Affected Product

| Product Name | Catalog | Lot No. | UDI-DI | Expiration Date |
| --- | --- | --- | --- | --- |
| BD BBL™ TB Stain Kit K | 212522 | 4191205 | (01)00382902125228 | 2025/04/30 |

Description of the Problem:

BD identified through customer complaints, that the TB CARBOLFUCHSIN KF 250ML component of the BD BBL™ TB Stain Kit K (catalog number 212522) was shipped with an expiration date of 30 Sep 2024 while the Certificate of Analysis (COA) indicated an expiration date of 30 Apr 2025.

Clinical Risk Statement:

If the expired product is used, there is a low potential risk of erroneous smear microscopy results, which could impact time to diagnosis and possibly delay initiation of treatment or lead to unnecessary treatment. No reports of patient harm or diagnostic inaccuracies have been received to date. The risk associated with use of the expired product is mitigated by good laboratory practice as well as the availability and use of additional diagnostic testing and clinical correlation in ascertaining a patient's true disease status.

Complaint & Adverse Event Statement:

To date, there have been no adverse events worldwide related to this event.

Actions for Clinical Users:

- Immediately inspect your inventory for the specific catalog and lot number listed to determine if the lot in your possession are impacted and immediately discontinue use of any remaining inventory of the affected product.

- Customers should verify product expiration dates on the reagent prior to use to prevent the use of expired materials. All other components of the BD BBL™ TB Stain Kit K are labeled with the correct expiration date of 30 Apr 2025.

- Ensure the contents of this notification are read and understood.

- Share and post this customer letter within your facility network and forward to any customers to whom you may have distributed the product to ensure awareness.

Action To be Taken by BD:

BD is investigating the root cause and will implement appropriate corrective actions to prevent recurrence of this issue.

Please Take the Following Actions:

- Please immediately discontinue the use and quarantine any of the unused inventory.

- Share this notification with all users within your facility network to ensure they are also aware of this medical device recall.

- If you purchased this product from a distributor, please contact your distributor for further return instructions and credit resolution.

- Complete the attached Customer Response Form and return it to your distributor and/or BD contact noted on the form indicating whether you have the affected lot so that BD may acknowledge your receipt of this notification.

- Indicate on the Customer Response Form the quantity for the affected lot identified at your facility and confirm that this product inventory was quarantined for return.

- Report any adverse health consequences experienced with the use of this product to BD.

BD is committed to advancing the world of health. Our primary objectives are patient and user safety and providing you with quality products. We apologize for any inconvenience this issue may have caused you and thank you in advance for helping us to resolve this matter as quickly and effectively as possible.

Yours Sincerely,

________________________________ Karena Han Quality Manager, Southeast Asia

CUSTOMER RESPONSE FORM

IDS-25-5245
BD BBL™ TB Stain Kit K

Please fill in the information below so that we may acknowledge your receipt of this notification.

Simply complete and return the completed form to SEA_Quality SEA_Quality@bd.com / distributor / local BD representative by 07 March 2025.

Please tick as appropriate.

|  | I have read and understood the attached notice & will share this Urgent Product Recall with all users within my facility. I shall clarify with the appointed distributor/ BD representative for clarification (s) to this Urgent Product Recall |
| --- | --- |
|  | We do not have affected product(s) in our inventory. |
|  | We have affected product(s) in inventory and will return to my distributor/local BD representative for destruction. |

| Product Description | Catalog No. | Lot No. | Quantity received (EA) | Quantity sold/used (EA) | Remaining Quantity (EA) in inventory to be *returned |
| --- | --- | --- | --- | --- | --- |

Completed by:

| Name: |  |
| --- | --- |
| Signature: |  |
| Date: |  |
| Facility / Address / Telephone Number: |  |