MEDICAL DEVICE PRODUCT ALERT

Re: Waters Instruments:

| Xevo TQ-S, IVD (PN: 186005453IVD) Xevo TQD IVD (PN: 186005833IVD) Xevo TQ-S Micro IVD (PN: 186007833IVD) Xevo TQ-XS, IVD (PN: 186008453IVD) | using Accessories: Kit, NeoLynx v4.2, IVD (PN: 667005682IVD) Kit, NeoLynx v4.1 IVD (PN: 667002653) Kit, IonLynx v4.2 IVD(PN: 667005757IVD) |
| --- | --- |

# 12Jan2021

Dear Customer:

Waters Corporation has initiated a Product Alert regarding our NeoLynx and IonLynx software applications.

After investigation a vulnerability has been found where the final sample of a batch may have missing functions when viewed within the NeoLynx or IonLynx processing browser.

The issue only presents when the last occurrence of the auto-refresh feature begins before all the data has processed. This is a race condition as the system is attempting multiple tasks that occasionally result in the refresh occurring before the processing has completed. A manual refresh or reopening of the file rectifies the issue.

Our records indicate that you have purchased NeoLynx and/or IonLynx software. To ensure continued reliability of our products, Waters has issued this product alert to make you aware of the issue referenced above and provide the following recommendations:

- Waters recommends the use of manual refresh at the end of an analysis.

- Waters also reminds customers to use QCs at the end of an analysis as directed by the Waters instructions for use and Clinical Laboratory Standards Institute (CLSI) guidance.

Please complete the attached Reply Verification Tracking Form and return it to Waters per the instructions within the form. Completion of this form provides evidence that Waters has notified you about this issue.

Feel free to contact me if you have any questions or concerns regarding this issue.

Sincerely,

Randy Koester

Sr. Director, Global Quality Assurance

Email:	  Randy_Koester@waters.com

Phone: 	  +1 508 482 2323

REPLY VERIFICATION TRACKING FORM

Medical Product Alert– Waters Instruments

Xevo TQ-S, IVD (PN: 186005453IVD), Xevo TQD IVD (PN: 186005833IVD), Xevo TQ-S Micro IVD (PN: 186007833IVD), Xevo TQ-XS, IVD (PN: 186008453IVD)

Using accessories: Kit, NeoLynx v4.2, IVD (PN: 667005682IVD), Kit, NeoLynx v4.1IVD (PN: 667002653), Kit, IonLynx v4.2 IVD (PN: 667005757IVD)

| Company Name: |  | Region: |  |
| --- | --- | --- | --- |
| Contact Person: |  | Email: |  |

Name of Product Affected

INSTRUCTIONS FOR COMPLETING THIS REPLY VERIFICATION TRACKING FORM:

This completed form is required for tracking purposes. Please check the applicable boxes below.  Your signature and date are required.

I acknowledge receipt of this letter

I have the affected product(s)

I DO NOT have the affected product(s)

Name (Print):  _________________________________________________________________________

Signature:  _______________________________________ Date: _________________

(DD/MM/YYYY)

Please return the completed form via email or fax to:

(email) waters_quality@waters.com

(fax) US: 508 482 2339 / Outside US: +1 508 482 2339