| No. | Date Received | Title of FCA | Affected Medical Device | MDA Reference Number | MDA Registration Number | Establishment Detail |
| --- | --- | --- | --- | --- | --- | --- |
| 1. | 03/06/2025 | ROTEM® sigma cartridges (Part No. 555501 & 555502) from certain lots may cause pressure errors during functional checks. | ROTEM® SIGMA | MDA/FCA/P1301-65308058-2025 | IVDC10235825-198676 | STRAITS SCIENTIFIC (M) SDN BHD straits@3s.com.my |
| 2. | 03/06/2025 | On ACL TOP systems, the HemosIL D-Dimer HS 500 reagent may carry over into the HemosIL Free Protein S assay, potentially affecting results. New mandatory parameters are provided to reduce this risk. | ACL TOP FAMILY | MDA/FCA/P1302-81066371-2025 | IVDC8262823-138692 |  |
| 3. | 08/06/2025 | Internal Standard Mix- Identified sporadic, but substantial interference ) which has led to an increase in signal intensity. | MASSCHROM® AMINO ACIDS AND ACYLCARNITINES FROM DRIED BLOOD (NON DERIVATISED | MDA/FCA/P1314-36051957-2025 | IVDC8881820-44301 | ANALISA RESOURCES (M) SDN BHD ltthang@analisa-scientific.com |
| 4. | 09/06/2025 | A software glitch on ORTHO VISION® and ORTHO VISION® Max analysers may prevent inventory of loaded samples or reagents due to timing issues with the loading station door. | ORTHO VISION ANALYZER SYSTEM | MDA/FCA/P1315-72182972-2025 | IVDD4853923-138156 | T T MEDICAL MANAGEMENT SDN BHD ar@ttmedical.com.my |
| 5. | 09/06/2025 | 2025-PD-CTAMI-004 Spectral CT with software version V5.0 or V5.2 Uncommanded motion may result in collision and additional software issues | SPECTRAL CT | MDA/FCA/P1315-33995668-2025 | GC6307722-96412 | PHILIPS MALAYSIA SDN BERHAD joshua.chan@philips.com |
| 6. | 10/06/2025 | 029 FSCA HC 25-01 ADVIA 2120 SETpoint and TESTpoint Controls ¦ Potential Low Bias for Patient Platelet Values and High Control Values for Platelets | ADVIA 2120I HEMATOLOGY SYSTEM | MDA/FCA/P1319-85603243-2025 | IVDB3655169817 | SIEMENS HEALTHCARE SDN. BHD. aili.siow@siemens-healthineers.com |
| 7. | 11/06/2025 | 2024-IGT-BST-015 Philips Azurion Systems RI.x and R2.x Software issue potentially resulting in loss of imaging (X-ray) functionality | AZURION 3 | MDA/FCA/P1318-19018758-2025 | GC99381544218 | PHILIPS MALAYSIA SDN BERHAD joshua.chan@philips.com |
| 8. |  |  | AZURION 5 | MDA/FCA/P1320-13417835-2025 | GC10785320-48186 |  |
| 9. |  |  | AZURION 7 | MDA/FCA/P1321-82432683-2025 | GC82521235717 |  |
| 10. |  |  | AZURION BIPLANE | MDA/FCA/P1322-27795446-2025 | GC59959884718 |  |
| 11. | 13/06/2025 | VITROS® 4600 Chemistry Systems and VITROS® 5600 and XT 7600 Integrated Systems May Incorrectly Allow Use of VITROS® Diluent Packs Identified as Expired | VITROS CHEMISTRY PRODUCTS ELECTROLYTES REAGENTS | MDA/FCA/P1325-38709250-2025 | IVDB7599223-142604 | T T MEDICAL MANAGEMENT SDN BHD ar@ttmedical.com.my |
| 12. | 18/06/2025 | 2024-IGT-IGTD-006 Visions PV.018 Digital IVUS Catheter Instructions for Use (IFU) Precaution Update | EAGLE EYE® PLATINUM DIGITAL IVUS CATHETER | MDA/FCA/P1326-56147634-2025 | GD772781341919 | PHILIPS MALAYSIA SDN BERHAD joshua.chan@philips.com |
| 13. | 19/06/2025 | FA1469 SmartSync Software Update to Eliminate Potential Erroneous Electrical Reset Message | CARELINK SMARTSYNC™ DEVICE MANAGER | MDA/FCA/P1327-14153342-2025 | GC2117624-178897 | MEDTRONIC MALAYSIA SDN BHD cherie.ang@medtronic.com |
| 14. | 19/06/2025 | FA1475 SmartSync Software Update to Align Abort Button Behavior Across Programmer Platforms | CARELINK SMARTSYNC™ DEVICE MANAGER | MDA/FCA/P1328-36402143-2025 | GC2117624-178897 |  |
| 15. | 24/06/2025 | FA1489 Nellcor™ Bedside SpO2 Patient Monitoring System | NELLCOR BEDSIDE SPO2 PATIENT MONITORING SYSTEM | MDA/FCA/P1330-92315695-2025 | GC64673492817 |  |
| 16. | 25/06/2025 | Frequency Inverter not programmed for Getinge Aquadis 56 | AQUADIS 56 | MDA/FCA/P1299-63848096-2025 | GB3393023-131060 | RBD HEALTHCARE SDN BHD sookpei@rbdhealthcare.com |