July 2, 2024

To: 	Sales Agents

Subject:	URGENT MEDICAL DEVICE CORRECTION

Affected Product: CPT® Hip System Femoral Stem 12/14 Neck Taper (see Attachment 2 for the Affected Product List)

Reason for Field Action Correction

Zimmer Biomet is voluntarily initiating a medical device field action correction for the CPT Hip System Femoral Stem 12/14 Neck Taper listed in Attachment 2 – Affected Product List due to an increased risk of postoperative periprosthetic femoral fracture (PFF).

The CPT Hip System Femoral Stem 12/14 Neck Taper is a polished taper-slip (PTS) style stem. Though PTS stems demonstrate low rates of aseptic loosening, contributing to high rates of survivorship, registry-based analyses have indicated that all PTS stems are associated with increased risk of PFF compared to composite beam  stems.3,4  Within the PTS stem category, stems manufactured from cobalt chromium alloy are at  increased risk for PFF in comparison to stainless steel.5  The CPT Hip System Femoral Stem 12/14 Neck Taper is manufactured from cobalt chromium alloy.

While PFF is a generally recognized risk associated with all total hip arthroplasty, the current instructions for use (IFU) for the CPT Hip System Femoral Stem 12/14 Neck Taper do not identify post-operative bone fracture as a risk. As a result, Zimmer Biomet is initiating a field action correction to ensure that orthopaedic surgeons are aware of this risk and that the CPT Hip System Femoral Stem 12/14 Neck Taper displayed a rate of PFF at approximately 1.4% in the United Kingdom, which is more than twice the risk of PFF relative to a similar PTS reference product that is stainless steel.6  The IFU is in the process of being updated to reflect the risk of PFF, with an anticipated release by the end of August 2024.

Risks

In the event that PFF occurs, which is often associated with a traumatic event, surgical intervention will likely be required. Surgical intervention typically involves either internal fixation of the fractured bone or the femoral stem implant is replaced with a new implant.  Users should consider potential risk factors for PFF that are available in published literature when evaluating its use during this time, such as those presented by Lamb et al.5 and ensure patients are informed appropriately on potential risks associated with the CPT Hip System Femoral Stem 12/14 Neck Taper. The potential modifiable risk factors include:

- Increasing age

- Intra-operative fracture

- Increasing stem offset

- Increasing head size

- Low viscosity bone cement

Additional Information on Affected Product

The CPT Hip System Femoral Stem 12/14 Neck Taper has been on the market for over 20 years. During this time, it achieved and has maintained the highest possible Orthopaedic Data Evaluation Panel (ODEP) rating of 15A*1, as well as provided 96.4% all-cause survivorship at 10 years in the latest United Kingdom National Joint Registry (UK NJR) annual report.2

Product Phase Out

Zimmer Biomet has also made the decision to phase out the CPT Hip System Femoral Step 12/14 Neck Taper, with a target date of December 2024 to stop sales in the Europe, Middle East and Africa (EMEA) region. However, no market exit date is currently confirmed for the Asia Pacific region.  The product has successful long-term survivorship and may continue to be implanted.

Zimmer Biomet recognizes the importance of a well-planned phase out and transition when changing hip stem brands. We will be taking actions to aid surgeons in transitioning to new hip stem brands. Once a market exit date for the Asia Pacific region is confirmed, we will communicate the dates for each country with the local authorities (as needed) as well as existing surgeon users.

Your Responsibilities:

- Review this notification and ensure that affected team members are aware of the contents.

- Complete Attachment 1 – Certificate of Acknowledgement and send to CorporateQuality.PostMarket@zimmerbiomet.com.

- Return the Additional Accounts form to CorporateQuality.PostMarket@zimmerbiomet.com.

  - Review the list of hospitals and/or surgeons included with the email notification sent to your facility, which includes a list of hospitals and/or surgeons that have already been notified of this field action.

  - Identify whether there are any additional hospitals and/or surgeons that have consumed the impacted devices within the past 12 months that Zimmer Biomet has not notified and list these accounts on the Additional Accounts form.  Please provide the form in Excel format.

  - If there are no additional accounts or surgeons to notify, please indicate that there are no additional accounts, or indicate “None” or “NA” on the form.

- Zimmer Biomet is in the process of updating the IFU to reflect the risk of PFF.  The latest IFU can be viewed electronically at labeling.zimmerbiomet.com and entering one of the part numbers from Attachment 2 – Affected Product List in the search field.

- Retain a copy of your Certificate of Acknowledgement form for your records in the event of a compliance audit of your facility.

- If you have further questions or concerns after reviewing this notice, please call customer service at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday.  Calls received outside of call center operating hours will receive a voicemail prompt or be transferred to an on-call representative in the event of an emergency.  Alternatively, your questions may be emailed to CorporateQuality.PostMarket@zimmerbiomet.com.

Other Information

This medical device field action was reported to the U.S. Food and Drug Administration and will be reported to other Competent Authorities, Notified Bodies, and Regulatory Authorities as required.

- Med Watch Reporting:  Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA’s Med Watch Adverse Event Reporting program either online, by mail, or by fax.

- Online:  www.fda.gov/medwatch/report.htm

- Call: 1-800-332-1088 to request a reporting form

- Mail:  Use postage paid, pre-addressed form FDA 3500, available at: www.fda.gov/MedWatch/getforms.htm

- Fax:  1-800-FDA-0178

Under 21 CFR 803, manufacturers are also required to report any serious injuries where a product has contributed or may have contributed to the event.  Please keep Zimmer Biomet informed of any adverse events associated with this product or any other Zimmer Biomet product by emailing product.experience@zimmerbiomet.com.

Please be aware that the names of user facilities notified are routinely provided to the Competent Authorities for audit purposes.  Your urgent cooperation is needed. The undersigned confirms that this notice has been delivered to the appropriate Regulatory Agencies.

Thank you for your assistance.  We regret any inconvenience caused by this field action.

Sincerely,

Sincerely,

_________________________________

Stephanie Leppo

Quality Associate Director

References

1Orthopaedic Data Evaluation Panel. www.odep.org.uk/product/cpt-stem-cocr/

2National Joint Registry. 20th Annual Report. 2023. www.njrcentre.org.uk

3Palan J, Smith MC, Gregg P et al. The influence of cemented femoral stem choice on the incidence of revision for periprosthetic fracture after primary total hip arthroplasty. Bone Joint J. 2016;98-B:1347-54. doi: 10.1302/0301-620X.98B10. PMID: 27694588.

4Jain S, Lamb JN, Pandit H. Cemented femoral stem design and postoperative periprosthetic fracture risk following total hip arthroplasty. Bone Joint J. 2024 Jan 1;106-B(1):11-15. doi: 10.1302/0301-620X.106B1.BJJ-2023-0587.R1. PMID: 38160687.

5Lamb JN, Jain S, King SW, West RM, Pandit HG. Risk Factors for Revision of Polished Taper-Slip Cemented Stems for Periprosthetic Femoral Fracture After Primary Total Hip Replacement: A Registry-Based Cohort Study from the National Joint Registry for England, Wales, Northern Ireland and the Isle of Man. J Bone Joint Surg Am. 2020 Sep 16;102(18):1600-1608. doi: 10.2106/JBJS.19.01242. PMID: 32604382.

6Pandit HG et al. Postoperative periprosthetic femoral fracture is the leading cause of major reoperation in the United Kingdom following primary total hip replacement: A study using national health data linked to National Joint Registry. Unpublished.

ATTACHMENT 1- Certificate of Acknowledgement

IMMEDIATE RESPONSE REQUIRED – TIME SENSITIVE ACTION NEEDED

Affected Product: CPT Hip System Femoral Stem 12/14 Neck Taper

Field Action Reference Number: ZFA 2024-00121

Sales Agent Company Name: _____________________________

Certificate of Acknowledgement:

By signing below, I acknowledge that I have received, read, and understand the contents of this correction communication. All required activities are complete or are being completed.

Printed Name: ____________________________       Signature: ______________________________

Title: ______________________________ Tel: (         ) ______________ Ext. _____ Date: __________

Note:  This form must be returned to Zimmer Biomet before this action is considered closed for your account.  It is important that you complete this form and email a copy to CorporateQuality.PostMarket@zimmerbiomet.com.

ATTACHMENT 2 - Affected Product List

| Item Number | GTIN Number | Item Description |
| --- | --- | --- |
| 00-8114-000-00 | 00889024145733 | Size 0 105 mm Stem Length Standard Offset |
| 00-8114-000-10 | 00889024145740 | Size 0 105 mm Stem Length Extended Offset |
| 00-8114-001-00 | 00889024145757 | Size 1 130 mm Stem Length Standard Offset |
| 00-8114-001-10 | 00889024145764 | Size 1 130 mm Stem Length Extended Offset |
| 00-8114-002-00 | 00889024145771 | Size 2 130 mm Stem Length Standard Offset |
| 00-8114-002-10 | 00889024145788 | Size 2 130 mm Stem Length Extended Offset |
| 00-8114-002-30 | 00889024145801 | Size 2 130 mm Stem Length Extra Extended Offset |
| 00-8114-003-00 | 00889024145818 | Size 3 130 mm Stem Length Standard Offset |
| 00-8114-003-10 | 00889024145825 | Size 3 130 mm Stem Length Extended Offset |
| 00-8114-003-30 | 00889024145849 | Size 3 130 mm Stem Length Extra Extended Offset |
| 00-8114-004-00 | 00889024145856 | Size 4 130 mm Stem Length Standard Offset |
| 00-8114-004-10 | 00889024145863 | Size 4 130 mm Stem Length Extended Offset |
| 00-8114-004-30 | 00889024145900 | Size 4 130 mm Stem Length Extra Extended Offset |
| 00-8114-005-00 | 00889024145917 | Size 5 130 mm Stem Length Standard Offset |
| 00-8114-005-10 | 00889024145924 | Size 5 130 mm Stem Length Extended Offset |
| 00-8114-005-30 | 00889024145931 | Size 5 130 mm Stem Length Extra Extended Offset |
| 00-8114-040-00 | 00889024145962 | Extra Small 85 mm Stem Length |
| 00-8114-050-00 | 00889024145979 | Small 95 mm Stem Length |
| 00-8114-002-18 | 00889024145795 | Size 2 180 mm Stem Length Standard Offset |
| 00-8114-003-18 | 00889024145832 | Size 3 180 mm Stem Length Extended Offset |
| 00-8114-004-20 | 00889024145870 | Size 4 200 mm Stem Length Extended Offset |
| 00-8114-004-23 | 00889024145887 | Size 4 230 mm Stem Length Extended Offset |
| 00-8114-004-26 | 00889024145894 | Size 4 260 mm Stem Length Extended Offset |
| 00-8114-012-18 | 00889024145948 | Size 2 180 mm Stem Length Valgus Neck |
| 00-8114-013-18 | 00889024145955 | Size 3 180 mm Stem Length Valgus Neck |